IMPORTANT NOTICE

# Maintaining Your PA200-Series IV

Maintenance Kit# CLN203 contains all the necessary components to keep your PA200 operating at peak efficiency for the next 6 months.

The kit contains:

1	part# GPC46 - Ceramic Generator Plate

2	part# FUS125 - 125 milliAmp High Voltage Fuses

6	part# WIPE9814 - Alcohol/Antistatic Wipe

The following is the recommended monthly program:

- TURN THE UNIT OFF AND UNPLUG IT FROM THE POWER SOURCE.

- Using a standard Phillips #2 screwdriver, remove the 4 screws securing the cabinet cover and lift the cover off.

- Gently slide the white ceramic generator plate straight up and out of the frame.

- Wipe any dust from the inside of the unit with a clean dry cloth.

- Using an alcohol wipe, carefully clean the generator frame from which you pulled the ceramic generator plate.  Be sure to thoroughly clean & wipe the inside area of the generator frame at the bottom, between the 2 upright posts, where the ceramic generator plate rests.

- The alcohol wipe can also be used to clean the generator plate and to remove any difficult residues, from other areas, that were not removed with the dry cloth.  The alcohol will evaporate quickly.

- Gently replace the white ceramic generator plate by sliding it straight down to the base of the generator frame.  Be sure that the ceramic generator plate slides down the 'slots' at each end and is resting on the base of the generator frame.

NOTE: Replace the old generator plate with a new generator plate at each 6th month cleaning interval.

- Replace the 4 cover screws.

- There are 2 fuse holders on the back of the machine.  The smaller fuse holder contains the 125 milliAmp High Voltage fuse.  If this 125 milliAmp fuse is blown (NO OZONE) replace it with one of the spares supplied with this kit.  This fuse will blow if the ceramic generator plate is broken or if the generator frame becomes contaminated.  Keep your spares for future use.

- The unit is ready to be put back into service.

FOR ADDITIONAL MAINTENANCE INFORMATION CONTACT YOUR DISTRIBUTOR,

OR ZONTEC CUSTOMER SERVICE AT:

1-727-341-0547   FAX: 1-727-302-9816

Available Maintenance Supplies	Part Number

Alcohol Wipes (12 per pack)		WIPE9814-12PK

Fuse .125 Amp Slo-Blo (6 per pack)	FUS125-6PK

Generator Plate (2 per pack)		GPC46-2PK

ACC\203.BOX.0598